# Mouse KLK1/Kallikrein 1 ELISA Kit

## CATALOG NO: IMSKLK1KT

## LOT NO: Sample

## INTENDED USE

For the quantitation of Mouse Klk1 concentrations in cell culture supernatants, cell lysates, serum and plasma (heparin, EDTA).

## BACKGROUND

Repeat steps a-b 2 additional times.  
  
Discard the wash buffer in the wells into an appropriate waste receptacle. Then, invert the plate on the benchtop onto a paper towel and tap the plate to gently blot any remaining liquid.  
  
Add 100 µl of the prepared 1x Avidin-Biotin-Peroxidase Complex into each well. Cover with the plate sealer provided and incubate for 40 minutes at RT (or 30 minutes at 37°C).  
  
Wash the plate 5 times with the 1x wash buffer:  
  
Discard the liquid in the wells into an appropriate waste receptacle. Then, invert the plate on the benchtop onto a paper towel and tap the plate to gently blot any remaining liquid. It is recommended that the wells are not allowed to completely dry at any time.  
  
Add 300 µl of the 1x wash buffer to each assay well. (For cleaner background incubate for 60 seconds between each wash).  
  
Repeat steps a-b 4 additional times.

## PRINCIPLE OF THE ASSAY

The Innovative Research Mouse Klk1 Pre-Coated ELISA (Enzyme-Linked Immunosorbent Assay) kit is a solid-phase immunoassay specially designed to measure Mouse Klk1 with a 96-well strip plate that is pre-coated with antibody specific for Klk1. The detection antibody is a biotinylated antibody specific for Klk1. The capture antibody is monoclonal antibody from rat and the detection antibody is polyclonal antibody from goat. The kit includes Mouse Klk1 protein as standards.

## OVERVIEW

\*The sensitivity or the minimum detectable dose (MDD) is the lower limit of the target protein that can be detected by the kit. It is determined by adding two standard deviations to the mean O.D. value of twenty (20) blank wells and calculating the corresponding concentration.

## TECHNICAL DETAILS

Product Name: Mouse KLK1/Kallikrein 1 ELISA Kit   
  
Reactive Species: Mouse  
  
Size: 96 wells/kit, with removable strips.  
  
Description: Mouse KLK1 / Kallikrein 1 / Tissue kallikrein ELISA Kit (96 Tests). Quantitate Mouse Klk1 in cell culture supernatants, cell lysates, serum and plasma (heparin, EDTA). Sensitivity: 12pg/ml. The brand indicates this is a premium quality ELISA kit. Each kit delivers precise quantification, high sensitivity, and excellent reproducibility. Only our most reliable and effective kits qualify as , guaranteeing top-tier results for your assays.  
  
Sensitivity\*: 12 pg/ml  
  
Detection Range: 62.5 pg/ml - 4,000 pg/ml  
  
Storage Instructions: Store at 4°C for 6 months, at -20°C for 12 months. Avoid multiple freeze-thaw cycles (Ships with gel ice, can store for up to 3 days in room temperature. Freeze upon receiving.)  
  
Uniprot ID: P15947  
  
Capture/Detection Antibodies: The capture antibody is monoclonal antibody from rat and the detection antibody is polyclonal antibody from goat.  
  
Specificity: Natural and recombinant Mouse Klk1  
  
Standard Protein: Expression system for standard: NS0; Immunogen sequence: I25-D261  
  
Cross-reactivity: This kit is for the detection of Mouse Klk1. No significant cross-reactivity or interference between Klk1 and its analogs was observed. This claim is limited by existing techniques; therefore, cross- reactivity may exist with untested analogs.

## PREPARATIONS BEFORE ASSAY

Please read the following instructions before starting the experiment.  
  
Read this manual in its entirety in order to minimize the chance of error.  
  
Confirm that you have the appropriate non-supplied equipment available.  
  
Confirm that the species, target antigen, and sensitivity of this kit are appropriate for your intended application.  
  
Confirm that your samples have been prepared appropriately based upon recommendations (see Sample Preparation) and that you have sufficient sample volume for use in the assay.  
  
When first using a kit, appropriate validation steps should be taken before using valuable samples. Confirm that the kit adequately detects the target antigen in your intended sample type(s) by running control samples.  
  
If the concentration of target antigen within your samples is unknown, a preliminary experiment should be run using a control sample to determine the optimal sample dilution (see Sample Preparation).  
  
To inspect the validity of experiment operation and the appropriateness of sample dilution proportion, a pilot experiment using standards and a small number of samples is recommended.  
  
Before using the kit, spin tubes to bring down all components to the bottom of the tubes.  
  
Don’t let the 96-well plate dry out since this will inactivate active components on the plate.  
  
Don’t reuse tips and tubes to avoid cross-contamination.  
  
Avoid using the reagents from different batches together.  
  
The kit should not be used beyond the expiration date on the kit label. Any variation in diluent, operator, pipetting technique, washing technique, incubation time or temperature, and kit age can cause variation in binding. Variations in sample collection, processing, and storage may cause sample value differences.

## KIT COMPONENTS/MATERIALS PROVIDED

|  |  |
| --- | --- |
| **Component** | **Quantity** |

## MATERIALS REQUIRED BUT NOT PROVIDED

• Microplate reader capable of reading absorbance at 450 nm. Incubator.  
• Automated plate washer (optional)  
• Pipettes and pipette tips capable of precisely dispensing 0.5 µl through 1 ml volumes of aqueous solutions. Multichannel pipettes are recommended for a large numbers of samples.

## REAGENT PREPARATION

Bring all reagents to room temperature before use.  
   
 Wash Buffer: Dilute Wash Buffer (25X) with distilled water. For example, if preparing 500 ml of Wash Buffer, dilute 20 ml of Wash Buffer (25X) into 480 ml of distilled water.  
   
 Standard: Reconstitute the standard with standard diluent according to the label instructions. This reconstitution produces a stock solution. Let the standard stand for a minimum of 15 minutes with gentle agitation prior to making dilutions.  
   
 Detection Reagent A and B: Dilute to the working concentration using Assay Diluent A and B, respectively.

## SAMPLE PREPARATION

## DILUTION OF STANDARD

1. Label 7 tubes, one for each standard: 4000 pg/ml, 2000 pg/ml, 1000 pg/ml, 500 pg/ml, 250 pg/ml, 125 pg/ml, and 62.5 pg/ml.  
 2. Pipette 300 µl of the Sample Diluent into each tube.  
 3. Pipette 300 µl of the reconstituted standard into the first tube and mix to create the 4000 pg/ml standard.  
 4. Pipette 300 µl from the 4000 pg/ml tube into the second tube and mix to create the 2000 pg/ml standard.  
 5. Continue this process for the remaining tubes.  
 6. The Sample Diluent serves as the zero standard (0 pg/ml).

## TYPICAL DATA / STANDARD CURVE

This standard curve is provided for demonstration only. A standard curve should be generated for each set of samples assayed.

|  |  |
| --- | --- |
| **Concentration (pg/ml)** | **O.D.** |

[{'concentration': '0.0', 'od\_value': '0.061'}]

## INTRA/INTER-ASSAY VARIABILITY

- Intra-Assay Precision: Three samples of known concentration were tested on one plate to assess intra-assay precision.

- Inter-Assay Precision: Three samples of known concentration were tested in separate assays to assess inter-assay precision.

## ASSAY PROTOCOL

1. It is recommended that all reagents and materials be equilibrated to room temperature (18-25°C) prior to the experiment (see Preparation Before The Experiment, if you have missed this information). Prepare all reagents and working standards as directed previously. Remove excess microplate strips from the plate frame and seal and store them in the original packaging. Add 100 µl of the standard, samples, or control per well. Add 100 µl of the Sample Diluent into the zero well. At least two replicates of each standard, sample, or control is recommended. Cover with the plate sealer provided and incubate for 120 minutes at room temperature (or 90 min. at 37 °C). Remove the cover and discard the liquid in the wells into an appropriate waste receptacle. Invert the plate on the benchtop onto a paper towel and tap the plate to gently blot any remaining liquid. It is recommended that the wells are not allowed to completely dry at any time. Add 100 µl of the prepared 1x Biotinylated Anti-Mouse Klk1 antibody to each well. Cover with a plate sealer and incubate for 90 minutes at room temperature (or 60 minutes at 37°C). Wash the plate 3 times with the 1x wash buffer: Discard the liquid in the wells into an appropriate waste receptacle. Then, invert the plate on the benchtop onto a paper towel and tap the plate to gently blot any remaining liquid. It is recommended that the wells are not allowed to completely dry at any time. Add 300 µl of the 1x wash buffer to each assay well. (For cleaner background incubate for 60 seconds between each wash). Repeat steps a-b 2 additional times. Discard the wash buffer in the wells into an appropriate waste receptacle. Then, invert the plate on the benchtop onto a paper towel and tap the plate to gently blot any remaining liquid. Add 100 µl of the prepared 1x Avidin-Biotin-Peroxidase Complex into each well. Cover with the plate sealer provided and incubate for 40 minutes at RT (or 30 minutes at 37°C). Wash the plate 5 times with the 1x wash buffer: Discard the liquid in the wells into an appropriate waste receptacle. Then, invert the plate on the benchtop onto a paper towel and tap the plate to gently blot any remaining liquid. It is recommended that the wells are not allowed to completely dry at any time. Add 300 µl of the 1x wash buffer to each assay well. (For cleaner background incubate for 60 seconds between each wash). Repeat steps a-b 4 additional times. Discard the wash buffer in the wells into an appropriate waste receptacle. Then, invert the plate on the benchtop onto a paper towel and tap the plate to gently blot any remaining liquid. Add 90 µl of Color Developing Reagent to each well. Cover with the plate sealer provided and incubate in the dark for 30 minutes at RT (or 15-25 minutes at 37°C). (The optimal incubation time must be empirically determined. A guideline to look for is blue shading the top four standard wells, while the remaining standards remain clear.) Add 100 µl of Stop Solution to each well. The color should immediately change to yellow. Within 30 minutes of stopping the reaction, the O.D. absorbance should be read with a microplate reader at 450nm. Assay Protocol Notes Solutions: To avoid cross-contamination, change pipette tips between additions of each standard, between sample additions, and between reagent additions. Also, use separate reservoirs for each reagent. Applying Solutions: All solutions should be added to the bottom of the ELISA plate well. Avoid touching the inside wall of the well. Avoid foaming when possible. Assay Timing: The interval between adding samples to the first and last wells should be minimized. Delays will increase the incubation time differential between wells, which will significantly affect the experimental accuracy and repeatability. For each step in the procedure, total dispensing time for addition of reagents or samples should not exceed 10 minutes.

## DATA ANALYSIS

Subtract the average zero standard O.D. reading. It is recommended that a standard curve be created using computer software to generate a four-parameter logistic (4-PL) curve-fit. A free program capable of generating a four-parameter logistic (4-PL) curve-fit can be found online at: www.myassays.com/four-parameter-logistic-curve.assay. Alternatively, plot the mean absorbance for each standard against the concentration. The measured concentration in the sample can be interpolated by using linear regression of each average relative O.D. against the standard curve generated using curve fitting software. This will generate an adequate but less precise fit of the data. For diluted samples, the concentration reading from the standard curve must be multiplied by the dilution factor. Background on Klk1  
  
Kallikrein-1, also known as tissue kallikrein, is a protein that in humans is encoded by the KLK1 gene. This serine protease generates Lys-bradykinin by specific proteolysis of kininogen-1. KLK1 is a member of the peptidase S1 family. Its gene is mapped to 19q13.3. In all, it has got 262-amino acids which contain a putative signal peptide, followed by a short activating peptide and the protease domain. The protein is mainly found in kidney, pancreas, and salivary gland, showing a unique pattern of tissue-specific expression relative to other members of the family. KLK1 is implicated in carcinogenesis and some have potential as novel cancer and other disease biomarkers. 1 Publications Citing This Product  
  
PubMed ID: 10.1186/s12014-021-09335-9, Proteomics and functional study reveal kallikrein-6 enhances communicating hydrocephalus Visit bosterbio.com/mouse-klk1-picokine-trade-elisa-kit-ek1586-innovative research.html to see all 1 publications.

## DISCLAIMER

This material is sold for in-vitro use only in manufacturing and research. This material is not suitable for human use. It is the responsibility of the user to undertake sufficient verification and testing to determine the suitability of each product's application. The statements herein are offered for informational purposes only and are intended to be used solely for your consideration, investigation and verification.